## Slide 1: Sealing and Stacking
Sealing
Except for some very basic array configurations, all microfluidic devices require a top cover to create enclosed structures, (e.g., channels, reservoirs). This can be achieved by sealing the top side of the substrate carrying the fluidic structures with a foil, a cover plate or by stacking several microstructured fluidic plates.
Sealing with a thin, and often optically transparent foil is a cost-effective procedure and allows easy access to the fluid, for example when optical methods are used as detection methods
Special materials may be selected for the foil such as foils with high diffusion coefficient for gasses thereby allowing oxygen supply to cells in the chips.

## Slide 2
Stacking of several microstructured plates is of advantage for more complex fluidic devices because it is extending integration into the vertical direction.
With stacking multilayer fluidic interconnections can be created and many fluidic devices are much easier to build when vertical integration is used.

## Slide 4
One example is that of micropumps, where the pumping chamber and valve seats may be on one plate, the membrane and the valve lids on a second and the driver and input and output on other plates.

## Slide 6: Methods
Three types of materials are common for microfluidic and Lab-on-a-Chip devices:
Silicon, glass, polymer materials
Silicon
Silicon is the dominant material in microelectronics, silicon has also been the dominant material used in nonelectronic applications and also in microfluidics.
Materials for the Manufacture of Microfluidic Components

## Slide 7
Silicon Advantages :
Simple generation of an inert surface (SiO2) by oxidation, high-temperature stability, high chemical resistance to organic solvents and acids, well-established bonding processes, an extensive knowledge about coating,…
Disadvantages:
Non-ideal surface for many biochemical applications,
High price for material and processing.

## Slide 8
Glass
Borosilicate types of glass are often used.
Boroflat or Pyrex 7740 have thermal expansion coefficients which are matched to that of silicon and are used together with silicon in stacked arrangements, e.g., as transparent cover plates.
These glasses can be bonded to silicon by anodic bonding without the need of a bonding material.

## Slide 9
Anodic bonding is a wafer bonding process to seal glass to either silicon or metal without introducing an intermediate layer; it is commonly used to seal glass to silicon wafers in electronics and microfluidics.
This bonding technique, also known as field assisted bonding or electrostatic sealing, is mostly used for connecting silicon/glass and metal/glass through electric fields.

## Slide 10
Glass Advantages: high chemical resistance, excellent thermal and mechanical stability, and optical transparency. It is well-suited as a surface for biological and chemical reactions.
Glass as also well-suited for capillary electrophoresis.
Glass as a base material is less expensive than silicon,

## Slide 11
Disadvantage:
High substrate and processing costs, and compared to other materials a high surface roughness of structures.
Polymers
The main benefit of polymer materials is based on simple and cost effective replication method such as injection molding or hot decoration, because this allows the manufacture of all microstructures of the device in one manufacturing cycle.

## Slide 12
These replication methods require the manufacture of a master structure which is used as a tool in the replication step.
As the manufacturing cost of the mastering is considerable these methods only make sense for high-volume applications.
Disadvantages :
Reduced thermal stability,
Reduced stability against organic solvents, acids and bases.

## Slide 13: Fluidic Structures
The most basic fluidic structures to built microfluidic devices are microchannels.
These channels provide the fluidic interconnection network between the fluidic elements of the device, but may have additional functions like the channels in capillary electrophoresis and other separation techniques.
Various shapes for the channel cross-sections are used, including rectangular, V-shaped and round.
The shape of the cross section may be determined by the fabrication method

## Slide 14
Important parameters of the microchannels include surface roughness and aspect ratio of the structure, which is defined as the ratio of depth to width in the case of channel.
High-aspect ratio channels have a high surface to volume ratio and consume less floor space on the microfluidic chip.
Channel widths commonly vary between the mm to μm range, aspect ratio up to 10 are used.

## Slide 15
One microfluidic device may carry channels of different widths and aspect ratios for different purposes.
For example auxiliary channels are used in capillary devices with a much smaller diameter than the fluid channels, to allow the air to exit from the device when it is filled with fluid by capillary forces.

## Slide 16
Other important structures are reaction/detection chambers, and sample and waste reservoirs.
These are larger, well-type structures, with dimensions often in the millimeter range, designed to hold the correct amount of fluid.
In the case of reaction chambers, it is often advantageous to generate a high surface to volume ratio. This can be achieved using auxiliary structures by folding up a channel in a meander-like form, or using a porous, nanostructured surface.

## Slide 17
One important case of microreaction chambers is that of microcompartments used in array-type microfluidic devices for parallel processing such as DNA chips and nanotiter plates.
These are designed to hold fluid volumes in the order of 10nL to several hundred of nL.
In the case of nanotiterplates the bottom compartments may be a thin membrane so that optical detection techniques can be applied through the membrane from the bottom side.

## Slide 18: Fabrication Methods
The fluidic structures are fabricated using standard methods of microfabrication.
These microfabrication methods are also used to manufacture the master tool for the microreplication of polymers. The master is usually a metal tool.
Practical marketable fluidic devices are generally multilevel structures. This means that a device will not carry structures of one common structural depth only, but will have channels, wells, and reservoirs with a variety of structural depths.

## Slide 19
This cannot usually be achieved with one single fabrication step not using just one fabrication technology, and in practice a combination of different fabrication technologies, each suitable for the generation of structures of the desired size, shape and precision will have to be applied.

## Slide 20: Surface modification
Modifications of the surface of the device are essential for the designed functionality of microfluidic device in (nano)biotechnology.
Often, these modifications are to be achieved locally, and therefore different areas of the device will require different modifications.
These modifications must be achieved on all surfaces, including the sidewalls of high-aspect ratio microstructures, for example in deep channels.

## Slide 21
Modification Types:
Modification of wetting characteristics (hydrophobic/hydrophilic)
Materials such as glass, Si and SiO2 have many OH-groups on their surface, and this causes hydrophilic behavior.
Increased biocompatibility, electroosmotic flow
Immobilizing the reagents, DNA, etc.

## Slide 22: The objectives for modifications of the fluidic device surface:
-To carry out chemical reactions or detection mechanisms
- To provide a proper surface for immobilization
- Increasing the surface area for catalytic reactions
- Fastening sieving matrices or stationary phases for separation devices.

## Slide 23
Polymer surfaces are hydrophobic in most cases.
Hydrophilic surfaces may be easily generated using O2 plasma treatment, but such surfaces are stable only for a few days.
More stable surface modifications are obtained by plasma polymerization of layers involving OH-groups at the surface.
The wetting characteristics of the surface may also be modified by a nanostructured surface.

## Slide 24
This principle of nanobiotechnology is found in the nature, for example in the circular structure of leaf surfaces and in the fractal surfaces.
Such water-repelling surfaces have self-cleaning properties, as particles on nanostructured hydrophobic surfaces are more readily wetted and wash away.

## Slide 25
Large surface areas are required for both catalytic reactions and separation assays and this may be achieved by coating microfluidic chips with a porous material.
In the case of silicon, porous silicon with pore sizes in the nanometer to micrometer range may be generated.
Another important surface functionalization is the binding of specific molecules to designed areas of the chip.
Such applications include DNA-, proteomics-, cell-, and tissue chips.

## Slide 26
Generally using various surface chemistries, linkers for such molecules must be provided in designated areas, while the remaining surface should be nonbinding.
Methods to immobilize the specific molecules include adsorption, crosslinking, covalent binding, microencapsulation and entrapment.
A thin , sputtered gold film can be used to immobilize a dense molecular film of thiols, providing a high density of alkyl groups as binding sites for surface reactions.

## Slide 27: Detection mechanisms
For most microfluidic applications detection devices are not integrated in to the fluidic chip but form part of a separate handling and detection system.
In this way the system can be re-used for the evaluation of a large number of chips.
One problem associated with detection in microfluidic devices is the small sample volume. For example due to the small dimensions of the system the optical pathlength for absorption measurments is also likely to be very small.
Commonly used detection methods include absorption (UV, optical, IR) Flu, electrochemical ,…

## Slide 29: Commercial Systems: Glucose detection
Roche; AccuCheck
Glucose in blood
Lifescan; SureStep Flexx
Glucose monitor

## Slide 30: Commercial Systems: Blood Analyzer
Abbott; i-STAT
Siemens; Quicklab
CARESIDE
Analyzer

## Slide 31: Commercial Systems
Affymetrix
Product: GenChip
DNA high density
array

## Slide 32: Commercial Systems
Advalytix
1 µl PCR slides for forensic

## Slide 33: Commercial Systems
Nanostream; Velove microparallel LC
(micro parallel liquid chromatography)
Tecan; LabCD
(development stop 7.05)
Gyros; Gyrolab